PACS Server Manual

Picture Archiving and Communication System

Basic Features:

- Connects to DICOM Compliant Modalities.: CT / MRI / Ultrasound / Digital X-Rays / Nuclear / Eye-Laser / Etc

- Send Email and SMS (w/ MEDSMS feature) of Image and Interpretation

- Modality Worklist Server – Machine retrieves patient information from server based on HIS Request

- Print past transactions to DICOM film printer

- Burn to CD

- Linked to Medsys Radiology Module

- Images are stored in DICOM (Digital Imaging and Communications in Medicine) format

- PACS Acquire (optional) supports video/image recording of non-DICOM machines, such as Endoscopy

Server

- MedsysDICOMserver.exe must be running at the PACS Server

DBServer.txt contains the connection settings to the database server

Example contents:

PROVIDER=SQLOLEDB;Server=medsys;uid=sa;pwd=pass;database=DICOM

- Network port 104 and File and Printer sharing should be enabled, a folder d:\images should be shared, this is where images will be stored

Modality Worklist (optional)

- Installed also on server, machines read from the Worklist to retrieve patient information from the HIS reqeuests.

- Machines should support Worklist, which is an optional component (contact your supplier for details); the IP of the server must be entered in the machine. AE Ttitle is MedsysPACSServer

- Uses port 105

Modality Machine/Instrument setup

- Enter the IP address of the PACS server, AE Title could be set to MedsysPACSServer, port is 104

Installation Configuration

| PACS Server IP |  |
| --- | --- |
| PACS Server computer name |  |
| PACS Server Admin username & password |  |
| PACS Server client username and password |  |
| Connected Modalities |  |
| PACS Acquire Modalities |  |
| Network Printer IP |  |
| Worklist enabled modalities |  |
| Reading workstations IP |  |
| DocModNet view enabled |  |
| PACS storage capacity/backup |  |
| SQL UserName and Password |  |
| Public IP address |  |
| Notes |  |

Client Setup:

- User login must be first created in the user manager with password Department code=23

- Under Options, enter the computer name or IP  of the PACS server, ex. \\pacsserver

- If the image retrieval mode is fileshare, make user \\pacsserver\images can be accessed from the workstation/terminal

Installation

Required Files:

Comdlg32.ocx

Msflxgd.ocx

Mshflxgd.ocx

tabctl32.ocx

Mscomctl.ocx

Mscomctlb.ocx

Mscomct2.ocx

Crviewer.dll

Dicomobjects.ocx

MSMAPI32.OCX

Sapi.dll

Scrrun.dll

Craxdrt.dll

Nerocom.dll

Prjmedsyslic.dll

> Copy the files locally (e.g. c:\medsys\pacs\ddl) and Run setup.bat to register the files

Database Files:

Build_File

Password

Patient Data

PACS

DICOM

Copy ntwdblib from crystal to \windows\system32

Options:

UsePACSDBExamcodes: unchecked

Firewall ports:

104 – DICOM Communication

105 – DICOM Worklist Communication

1433 – SQL

3389 – Remote Desktop

Server Setup:

Share folder \IMAGES\

Connecting:

- Under password, type “ Medsys” (space medsys) and click on Server…Type the server name and userid/pw then click Test Connection.

Hardware Recommendations:

Reading Terminals Display:

- 19” or bigger LCD, 1280x1024 / 1600x1200 or more resolution.  Two LCD panels per terminal, for better image comparison.  Medical-grade panels are good but come with a higher cost.  Swivel panel.

- Video card with DVI-out and dual view capability.

Viewing terminals need not be upgraded.

PACS Server:

- Expandable RAID storage (RAID 10). 1 Terabyte initial capacity is good.

- Gigabit LAN, connected to database server.

Average File Size

- CT/MR - ~200kb per image (total size depends on the machine's number of slices).  Estimated average of ~25Meg per exam

- US - ~450kb per image (~3Meg per exam)

- CR (scanned image) - (~5Meg per image), 1 to 2 images per exam.

- Endoscopy (PACS Aquire) - AVI video (~300Meg)

PACS Client Manual

Client

- To enable image smoothing, under Options (Viewer tab), set magnification to B-spline or Cubic spline.  Some doctors prefer smoothing; some prefer the default pixel value

|  | To set the default modality, first close the View exam window, check one modality from the PACS Menu screen, and then click View Exam. |
| --- | --- |

- To open an image, select Open Exam from toolbar – Click on Search then double-click on Patient Name

- If you wish to export the image to BMP or JPG click File | Export to non-DICOM file from menu

- To open images from other series, double click the series description on the Series Description list

- To change the level/width of X-ray images, first uncheck the VOILUT checkbox.  The VOILUT checkbox sets the optimal viewing configuration.   If VOILUT is not used, click “def” to set the image to the default level/width.

- Play Images – circulates through the opened images, click Play Images again to stop the process

- Open slice window – could be configured to open on the secondary monitor, mainly used for multi-slice CT

- Click on Previous Result to open past images.

- Click on Stretch to fit and uncheck the box to see the image in the default resolution.

|  | Mouse Functions: Arrow – Default Hand – Moves image up/down/left/right. Left click on the image and hold left mouse button down, then drag the mouse in the desired direction. Wand – Measure A – Text Ma – Magnifiication Sm – Smoothing option (selected image) W2/W1 – adjust Level/Width or Brightness/Contrast. To use, left-click on center of image, hold mouse down then move mouse up or down to increase/decrease level, and left and right to change width. Line – draws a line Pen – Draws free form Square,Circle – draws respective object Po – Protractor Ci – plays Cine movie for Cine images |
| --- | --- |

|  | HIS Link When Worklist is not used: On the machines, enter the HIS Request Number in the accession number field, if the correct request/accession number is not entered, this could be adjusted using the HIS Link. Search for the corresponding request buy entering the Request Date, and Last name. Once a request has been highlighted, click on Save Accession Num to update the image accession number to the selected request number. |
| --- | --- |

MEDSYS PACS

Quick Guide

- Double-click MEDSYS PACS icon on the desktop.

- Enter your MEDSYS user name and password (given by EDP).

- Click on the name of the patient to open the images.

  - Or type the Last Name of the patient and click Search

  - If name does not appear, make sure the modality (at the top) is checked.

- If multiple images are displayed, double-click the image to fit it to the whole screen.

- Select series - If there are several series in the study, move mouse to the left side of the screen to select the series.

- Tools - move mouse to the bottom of the screen to access the tools.

Basic Tool Functions: (Can be accessed also using Right Click)

- Windowing - Click and hold left mouse button then drag Up/Down or Left/Right to adjust the brightness and contrast.

- Pan - Click and hold left mouse button then drag to move image around.

- Zoom - Click and hold left mouse button then drag to zoom image.  Alternatively, use the mouse wheel to zoom (no need to click on zoom).

- Calibrate 1cm – (for pixel images) Click and hold left mouse button then drag until 1cm is reached.  Click and hold left mouse button again to measure in cm.

- Right-click tools:

- Capture Part and Send – Makes a copy of the image with annotations and focused area then transmits to the PACS server.

- Previous Exam – If a previous exam exists, a blue previous exam button will appear.

- Left/Right Arrow Keys – Use the arrow keys to go the next/previous image.