PACS Manual

Basic Features:

- Connects to DICOM Compliant Modalities.: CT / MRI / Ultrasound / Digital X-Rays / Nuclear / Eye-Laser / Etc

- Send Email and SMS (w/ MEDSMS feature) of Image and Interpretation

- Modality Worklist Server – Machine retrieves patient information from server based on HIS Request

- Print past transactions to DICOM film printer

- Burn to CD

- Linked to Medsys Radiology Module

Installation

Required Files:

Comdlg32.ocx

Msflxgd.ocx

Mshflxgd.ocx

tabctl32.ocx

Mscomctl.ocx

Mscomctlb.ocx

Mscomct2.ocx

Crviewer.dll

Dicomobjects.ocx

MSMAPI32.OCX

Sapi.dll

Scrrun.dll

Craxdrt.dll

Nerocom.dll

Prjmedsyslic.dll

Database Files:

Build_File

Password

Patient Data

PACS

DICOM

Copy ntwdblib from crystal to \windows\system32

Setup Options:

Dicomserver (ip address of server): ex: \\192.168.8.8

-make sure user has rights to that folder….type the ip address in start | run.   ex: \\192.168.8.8

UsePACSDBExamcodes: unchecked

Firewall ports:

104 – DICOM Communication

105 – DICOM Worklist Communication

1433 – SQL

3389 – Remote Desktop

Server Setup:

Share folder \IMAGES\

Hardware Recommendations:

Reading Terminals Display:

- 19” or bigger LCD, 1280x1024 / 1600x1200 or more resolution.  Two LCD panels per terminal, for better image comparison.  Medical-grade panels are good but come with a higher cost.  Swivel panel.

- Video card with DVI-out and dual view capability.

Viewing terminals need not be upgraded.

PACS Server:

- Expandable RAID storage (RAID 10). 1 Terabyte initial capacity is good.

- Gigabit LAN, connected to database server.

Average File Size

- CT/MR - ~200kb per image (total size depends on the machine's number of slices).  Estimated average of ~25Meg per exam

- US - ~450kb per image (~3Meg per exam)

- CR (scanned image) - (~5Meg per image), 1 to 2 images per exam.

- Endoscopy (PACS Aquire) - AVI video (~300Meg)